Battery Voltage Measurement - ADC Calibration

The MCP3021 (12-bit) ADC was used to measure the battery voltage in the custom ESP32-H2 data logger hardware.

- Device Under Test (DUT): Custom PCB with ESP32-H2 and MCP3021 ADC

- Reference Instrument: Digital Multimeter (DMM, ±25mV at range (9.999V))

- Voltage Source: Bench DC power supply

- Measurement Range: 3300–4300 mV

Measurement Data:

Multiple voltages measured (Supplied from Bench power supply in range 3.3V to 4.3V) simultaneously by the DMM and ADC (uncalibrated firmware output).

Table 1.1: Reference Voltage and ADC computed Voltage in mV

| Reference Voltage (mV) [±25 mV] | ADC Measurement (mV) |
| --- | --- |
| 3302 | 3317 |
| 3401 | 3409 |
| 3501 | 3506 |
| 3601 | 3606 |
| 3701 | 3703 |
| 3804 | 3802 |
| 3903 | 3899 |
| 4001 | 3999 |
| 4102 | 4098 |
| 4201 | 4201 |
| 4304 | 4294 |

Linear Regression Model

A linear model was fitted to relate the ADC reading (input) to the true reference voltage,

VCalibrated=a ×VADC+b

Best-fit parameters:

- Slope (a): 1.0188

- Offset (b): –72.6815 mV

- Coefficient of determination (R²): 0.999917

Residual statistics

- Mean residual: ≈ 4.5×10-13mV → effectively 0 mV

- Standard deviation: 2.89 mV

- Maximum absolute residual: 6.3 mV (at 4201 mV point)

Note: -

After calibration, the ADC readings follow the model with max error ≈ ±6.3 mV, but the absolute accuracy is limited by the DMM’s ±25 mV.

Fitted model with uncertainty:

VCalibrated mV=1.0188 ×VADC(mV)-72.6815   ±25 mV